## Slide 1
PARcast: E_d(z, PAR) breadboard wiring
In-water profiler: top-down view on a prototyping breadboard
+
+
–
–
+
+
–
–
5
5
10
10
15
15
20
20
25
25
30
30
35
35
40
40
45
45
50
50
55
55
60
60
a
a
b
b
c
c
d
d
e
e
f
f
g
g
h
h
i
i
j
j
ADS1115
16-bit ADC
VDD
GND
SCL
SDA
ADDR
ALRT
A0
A1
A2
DS3231
Real-time clock
Vin
GND
SCL
SDA
SQW
32K
RST
MS5837
Pressure (0x76)
VDD
GND
SCL
SDA
Teensy 4.1
GND
3V3
19 (SCL0)
18 (SDA0)
Built-in
microSD slot
USB on right end
Apogee SQ-500
PAR sensor (0–40 mV, self-powered)
Serial 1559+: White / Black / Clear (grey)
White → A0
Black → A1
Clear (grey) → GND (shield)
Wire legend
Red — 3.3 V
Black — GND
Yellow — SCL (I2C clock)
Blue — SDA (I2C data)
White — SQ-500 signal +
Gray dashed — SQ-500 shield
Notes
• I2C bus shared by ADS1115, DS3231, MS5837
• Addresses: 0x48, 0x68, 0x76
• Teensy SDA0 = pin 18, SCL0 = pin 19
• SQ-500 differential: reads A0 − A1
• SQ-500 self-powered, no Vcc
• Underwater: apply 1.25× immersion factor
• MS5837 sealed via 1.8×0.8 mm O-ring
PARcast — github.com/LoriAzniveBerberian/PARcast

## Slide 2
E_d(z, PAR) connection table
Authoritative pin-by-pin reference for build and debug
From | Pin / wire | To | Pin / wire | Wire color | Purpose
Teensy 4.1 | 3V3 | Breadboard | Top + rail | Red | Main 3.3 V distribution
Teensy 4.1 | GND | Breadboard | Top – rail | Black | Main ground distribution
Top + rail | — | ADS1115 | VDD | Red | ADC power
Top – rail | — | ADS1115 | GND | Black | ADC ground
Top + rail | — | DS3231 | Vin | Red | RTC power
Top – rail | — | DS3231 | GND | Black | RTC ground
Top + rail | — | MS5837 | VDD | Red | Pressure sensor power
Top – rail | — | MS5837 | GND | Black | Pressure sensor ground
Teensy 4.1 | Pin 18 (SDA0) | ADS1115 | SDA | Blue | I2C data
Teensy 4.1 | Pin 18 (SDA0) | DS3231 | SDA | Blue | I2C data (shared bus)
Teensy 4.1 | Pin 18 (SDA0) | MS5837 | SDA | Blue | I2C data (shared bus)
Teensy 4.1 | Pin 19 (SCL0) | ADS1115 | SCL | Yellow | I2C clock
Teensy 4.1 | Pin 19 (SCL0) | DS3231 | SCL | Yellow | I2C clock (shared bus)
Teensy 4.1 | Pin 19 (SCL0) | MS5837 | SCL | Yellow | I2C clock (shared bus)
SQ-500 | White wire | ADS1115 | A0 | White | Sensor signal positive
SQ-500 | Black wire | ADS1115 | A1 | Black | Sensor signal negative
SQ-500 | Clear shield | Bottom – rail | — | Gray (dashed) | Shield to ground
Notes on bench testing
• Run I2C scanner sketch: 0x48 (ADS1115), 0x68 (DS3231), and 0x76 (MS5837) must all appear.
• If only some addresses show, check SDA/SCL jumpers and common ground.
• SQ-500 is self-powered: never connect it to a voltage rail.
• MS5837 PROM CRC is checked at startup — failure indicates wiring or sealing damage.
• Underwater PPFD = air PPFD × 1.25 (immersion correction for SQ-500-SS, S/N ≥ 2876).
PARcast — github.com/LoriAzniveBerberian/PARcast